RXPert® Single Filtered Balance Systems & RXPert® Double Filtered Balance Systems,

LABCONCO RXPert® Filtered Balance System Specification 

PART 1 - GENERAL 

1.01 Work Included: 

1.01.01 This specification covers the requirements for the purchase of RXPert® Filtered Balance Systems. 

1.01.02 Bench-mounted RXPert® Filtered Balance System in 2, 3, 4, 5 and 6 foot widths are covered by this specification. 

1.01.03 This specification sets the intent for quality, performance and appearance. 

1.02 Quality Assurance:

1.02.01 The manufacturer maintains a testing facility at their place of business for the performance testing of bench-mounted RXPert® Filtered Balance System. Both system and installation are in conformance to good construction practice and approved by the owner/user. The test facility as well as the manufacturing facility are available for owner/user inspection and its quality control procedures. All balance systems wired for 115 volts, 60 Hz meet or exceed all minimum requirements of UL Standard 3101-1 and CAN/CSA C22.2 No. 1010.1-92 and carry the ETL Testing Laboratories seal in the U.S. and Canada. 
1.03 References: 

1.03.01 The bench-mounted RXPert® Filtered Balance System conform to the following regulations and standards: 

UL -- Standard 3101-1/61010-1 (115 volt, 60 Hz models only)
CAN/CSA -- C22.2 No. 1010.1-92 (115 volt, 60 Hz models only) 
SEFA 1-2002
Modified ASHRAE 110-1995
ANSI Z9.5-1993
 
1.04 Submittals: 

1.04.01 Bench-mounted RXPert® Filtered Balance System specification sheets and product manuals will be submitted by the manufacturer upon request. The balance system supplier will submit shop drawings when necessary for clarification. 

1.04.02 A copy of the test results conducted prior to shipping to ensure proper operation will be shipped with each cabinet. 

1.05 Delivery and Storage: 

1.05.01 Bench-mounted RXPert® Filtered Balance Systems are delivered adequately protected from damage during shipment. 

1.06 Warranty 

1.06.01 Manufacturer's warranty against defects in material or workmanship on its balance systems for 1 year from date of installation or 2 years from date of purchase, whichever is sooner, and includes replacement of parts (except HEPA filters and lamps) and labor. 

PART 2 - PRODUCTS

2.01 Acceptable Manufacturer:

2.01.01 Labconco Corporation, 8811 Prospect Avenue, Kansas City, Missouri 64132, 
RXPert® Filtered Balance System, model numbers as described below:

2’ Foot Models: 3971200, 3971201, 3971202, 3971280, 3971281, 3971282

3’ Foot Models: 3971300, 3971301, 3971302, 3971380, 3971381, 3971382

4’ Foot Models: 3971400, 3971401, 3971402, 3971475, 3971476, 3971477, 3971480, 3971481,     3971482, 3971485, 3971486, 3971487, 3971490, 3971490, 3971491, 3971492, 3971495, 3971496, 3971497
5’ Foot Models: 3971575, 3971576, 3971577, 3971585, 3971586, 3971587, 3971590, 3971591, 3971592, 3971595, 3971596, 3971597

6’ Foot Models: 3971675, 3971676, 3971677, 3971685, 3971686, 3971687, 3971690, 3971691, 3971692, 3971695, 3971696, 3971697

2.02 Materials: 

2.02.01 Static-dissipating, epoxy-coated aluminum is used for the frame.

2.02.02 Static-dissipating, epoxy-coated steel is used for plenums and baffle.  No unreliable plastic to plastic bonding is used. 

2.02.03 Static-dissipating, tempered safety glass is used for side walls and front sash.

2.02.04 HEPA filters are a minimum of 99.99% efficient on all particles 0.3µm. HEPA filters are industry-standard size.

2.02.05 Fluorescent lighting provides 28 to 35 foot-candles on work surface.

2.02.06 Tempered safety glass sash and side panels are 1/4" thick tempered safety glass.

2.02.07 Pressure gauge monitors system pressure and is located on the front panel. 

2.02.08 Sparkless motorized impellers are constructed of backward curved centrifugal fans highly resistant to particles with maintenance-free ball bearings and run over 50,000 hours. Two and three-foot models use 1/6 HP impellers.  Four, five and six-foot models use 1/3 HP impellers.  Five and six-foot models have 2 each motorized impellers.  

2.02.09 Built-in impeller is mounted on vibration isolation supports.

2.02.10 Impeller has a 40% reserve so that speed may be adjusted as needed as filter loads.

2.02.11 Internal ductwork is epoxy-coated steel.

2.02.12 Control panel has separate rocker switches that control blowers and lights.

2.02.13 Upper diffuser screen is constructed of epoxy-coated aluminum.

2.02.14 Optional work surface is dished to contain spills and constructed of solid epoxy or stainless steel.

2.02.15 Optional adjustable height base stand (not included with balance system) is epoxy-coated steel and ADA compliant.

2.03 Fabrication: 

2.03.01 Overall exterior dimensional information on benchtop filtered balance systems is as described in the table below.

Model Number*   Description	Actual Width	Actual Depth    Actual Height

3971280	   2 Foot Model	   24.0"		   29.0"		49.5"
3971281	   2 Foot Model	   24.0"		   29.0"		49.5"                     		            3971282	   2 Foot Model	   24.0"		   29.0"		49.5"

3971380	   3 Foot Model	   36.0"		   29.0"		49.5"
3971381	   3 Foot Model	   36.0"		   29.0"		49.5"                               		  3971382	  3 Foot Model	   36.0"		   29.0"		49.5"

3971480	   4 Foot Model	   48.0"		   29.0"		49.5"
3971481	   4 Foot Model	   48.0"		   29.0"		49.5"                              		   3971482	   4 Foot Model	   48.0"		   29.0"		49.5"

3971485	   4 Foot Model	   48.0"		   29.0"		58.8"
3971486	   4 Foot Model	   48.0"		   29.0"		58.8"                         		        3971487	   4 Foot Model	   48.0"		   29.0"		58.8"

3971490	   4 Foot Model	   48.0"		   35.5"		58.8"
3971491	   4 Foot Model	   48.0"		   35.5"		58.8"
3971492	   4 Foot Model	   48.0"		   35.5"		58.8"

3971585	   5 Foot Model	   60.0"		29.0"		58.8"
3971586	   5 Foot Model	   60.0"		29.0"		58.8"                         		        3971587	   5 Foot Model	   60.0"		29.0"		58.8"

3971590	   5 Foot Model	   60.0"		35.5"		58.8"
3971591	   5 Foot Model	   60.0"		35.5"		58.8"                               		  3971592	   5 Foot Model	   60.0"		35.5"		58.8"

3971685	   6 Foot Model	   72.0"		29.0"		58.8"
3971686	   6 Foot Model	   72.0"		29.0"		58.8"                         		        3971687	   6 Foot Model	   72.0"		29.0"		58.8"

3971690	   6 Foot Model	   72.0"		35.5"		58.8"
3971691	   6 Foot Model	   72.0"		35.5"		58.8"                                		 3971692	   6 Foot Model	   72.0"		35.5"		58.8"

*Models ending in -1 include stainless steel worksurface. Models ending in -2 include epoxy worksurface.

2.03.02 Sash is angled approximately 10° on two, three, and four-foot models with a fixed opening of 9.4 inches and approximately 7° on four-foot (Tall), five-foot and six-foot models with a fixed opening of 12.0 inches for better visibility. Sash has no visibility-interfering protrusions. 

2.03.03 Sash is hinged to pivot upward and locks to a loading height of 16.5 inches and max loading height of 19.0”.  Four-foot (tall), five-foot and six-foot models have a sash that incorporates an external gas-assist lift and locks to a loading height of 25.75 inches and max loading height of 28.25”.

2.03.04 Access panel on Double Filtered System models allows for 100% of primary HEPA filter. Upper diffusion screen allows for 100% HEPA filter scan on secondary HEPA filter.

2.03.05 HEPA filter is located downstream from blower to allow for 100% filter scanning.

2.03.06 The filtered balance enclosure contains a true bag-in/bag-out filter disposal system for HEPA filter(s).

2.03.07 Clean-Sweep™ Air Foil allows air to sweep the work surface for maximum containment. Air foil is ergonomic to allow comfortable resting of arms.

2.03.08 Side-Entry Air foils are aerodynamic and located on both sides of sash opening to direct room air along sides of the filtered balance system.

2.03.09 Upper dilution air supply allows air into the filter balance system from above the work area which bathes the back of the sash with room air and directs concentrations away from the sash opening.

2.03.10 Upper containment sash foil allows room air to bleed in between the sash and the handle to reduce turbulence and direct concentrations away from the users breathing zone

2.03.11. Zone-perforated rear baffle contains three zoned sections of perforations to direct air in a non-turbulent laminar air stream from the sash opening

2.03.12 The internal blower has a variable speed controller to allow user defined face velocities

2.03.13 Two utility ports are located in either corner on the lower rear wall to allow for pass-through of electrical cords.

2.03.14 The positive pressure motorized impeller is surrounded by a negative pressure plenum so that if a leak should develop, the unfiltered air is captured and directed through the HEPA filter(s).

2.03.15 A waste chute on the left side of the enclosure is standard on five and six-foot, 35.5 inch deep models and optional on other models. The waste chute is 5.2 inches in diameter and requires a minimum bag size of 10 inches by 16 inches.

2.03.16 The included Guardian Airflow monitor will continuously monitor airflow with a green LED when airflow is above a selected point and will display a red LED and sound an audible alarm when airflows drops below a set velocity.



2.03.19 An 8.2-foot power cord connects to the filtered balance system by a suitable keyed connector (IEC 320 AC inlet connector or equivalent). Permanently attached cords with strain relief connectors are not used.

2.03.20 Optional trace odor carbon filters are available for work with nuisance odors that are organic, aldehyde, or ammine based on Double Filtered models that recirculate air back into the work environment.

2.03.21 Optional utility shelves constructed of epoxy coated aluminum hook into rear baffle perforation to keep items off of work surface.

2.03.22 Optional Canopy Connection Kit and Damper for thimble ducting Double Filtered models to the outside are available.

2.03.23 Optional adjustable-height base stand (not included with cabinet) provide a working height from 30 to 36". Dimensions are as follows:

Nominal Dimensions 
2 ft. wide: 24" w x 29" d x 27.5"-33.5" h 
3 ft. wide: 36" w x 29" d x 27.5"-33.5" h
4 ft. wide: 48" w x 29" d x 27.5"-33.5" h 
5 ft. wide: 60" w x 29" d x 27.5"-33.5" h
6 ft. wide: 72" w x 29" d x 27.5"-33.5" h 

PART 3 - EXECUTION

3.01 Inspection 

3.01.01 Carefully check the contents of the carton for damage that might have occurred in transit. 

3.02 Preparation

3.02.01 Verify equipment rough-in before proceeding with work. 

3.02.02 Coordinate with other trades for the proper and correct installation of electrical rough-in and for rough opening dimensions required for installation of the filtered balance system.

3.03 Installation 

3.03.01 Install according to manufacturer's instructions. 

3.03.02 Install according to standards required by authority having jurisdiction. 

3.03.03 Install equipment square and straight with no distortion and securely anchor as required. 

3.03.04 Touch up minor damaged surfaces caused by installation. Replace damaged components as needed.

3.04 Field Quality Control 

3.04.01 A qualified independent certifier should perform a HEPA filter leak test to verify filter(s) damage has not occurred during transit. 

3.05 Cleaning 

305.01 Clean equipment, casework, countertops and all other surfaces as recommended by the manufacturer, rendering all work in a new and unused appearance. 

3.05.02 Clean adjacent construction and surfaces, which may have been soiled in the course of installation of work in this section. 

3.06 Protection

3.06.01 Provide all necessary protective measures to prevent exposure of equipment and surfaces from exposure to other construction activity. 

3.06.02 Advise contractor of procedures and precautions for protection of material and installed equipment and casework from damage by work of other trades.